## Slide 1: FRAMELESS STEREOTACTIC LOCALIZATION AND GUIDANCE

## Slide 2
Roberts and Friets -1980s.
Theory
- Three points define a volume in geometric space
- If the same three points can be defined on a patient and on an image of that patient, then instruments can be navigated.

## Slide 3: Basic components
Computer workstation with neuronavigation software and computer monitor
Medical imaging input (either through a DICOM link or a portable digital data format)
Optical digitizer with infrared emitters and two infrared cameras
Reference frame (secured to a head clamp)
Registration stars (frames) for surgical instruments.
Passive infrared reflectors (aluminum impregnated glass spheres)
Notes: Put only Photograph

## Slide 4: Understanding the Basics
The principal concepts that underpin neuronavigation:
Pre-operative imaging: Sequential , Non-overlapping volumetric slices, MRI/CT slices of 1-3 mm
Accurate registration of the ‘image dataset’ with the ‘real time surgical space’.
A triangulation based system (analogous to Satellite navigation) to track operative instruments.
A dynamic referencing system to maintain the validity of registration during the operative procedure.

## Slide 5: Pre-operative Imaging
Slices are contiguous without overlaps or spaces, the summation of the slices creates a 3-D reconstruction of the brain
Software enables the imaging dataset to be viewed in multiple planes and as a 3-D reconstruction.
Image fusion software can now fuse digitized information from other modalities such as CT and PET with the MRI scan at this stage.
Notes: Photogaphs…conver into two slides

## Slide 6: Registration
Point or Surface alignment technique
Point technique : Fiducial or Anatomical landmark.
Anatomical landmarks identifiable on the imaging and in the surgical field (e.g. nasion, external auditory canal, orbital margins) can be co-aligned for registration.
Notes: Photographs

## Slide 7
Surface alignment: Co-align the surface of the patient’s head with the surface of the image dataset. A Laser beam is used within sight of the optical digitizer to delineate the periorbital and forehead regions of the scalp.

## Slide 8
The accuracy of registration must be confirmed with visual checks.
Known landmarks on dataset are identified in the surgical space using the default pointer tool (e.g. globes, orbits, EAM, tragus).

## Slide 9: Dynamic Referencing
An optical reference frame with passive infrared reflectors is secured to the Mayfield head clamp.
The optical digitizers recognize the fixed relationship between the reference frame and the patient’s head.

## Slide 10: Neuro-navigation errors
Image dataset acquisition
Overlapping or interspacing slices.
Variable slice thickness due to mechanical factors
Resolution errors and limitations
Errors associated with image fusion algorithms
Prevention
Careful planning on well maintained equipment

## Slide 11
Registration
Adverse skin mobility
Poor delineation of occipital region in surface mapping models
Fiducial movement
Prevention
Careful placement of Mayfield clamp
Combination of surface and point registration
Inclusion of the whole head during registration

## Slide 12
Intra-operative errors
Parallax type error when planning craniotomy
Head movement causing loss of registration
Brain shift
Small errors in trajectory angle are magnified at the tip of a biopsy needle

## Slide 13
Prevention
Careful positioning
Prudent use of mannitol
Plan flap perpendicular to skull
Ensure head secure in Mayfield clamp
Use a rigid biopsy needle
Use the largest registration star possible when tracking instruments

## Slide 14: Brain shift
Brain shift :
CSF drainage
use of diuretics
tumor resection which makes an error in the accuracy of the preoperative data.
Deeper structures, such as those at the skull base or along the falx shift less than superficial structures, improving the accuracy of stereotactic systems at those regions.

## Slide 15
Steps to prevent Brain shift
Displacement is often greatest in the direction of gravity
Minimizing diuretic use
CSF shunting should be avoided until after the primary approach has been made
defining the outer margins of a tumor before proceeding with debulking

## Slide 16: Fiducial versus nonfiducial
Fiducial versus nonfiducial neuronavigation assessment and consideration of accuracy. Pfisterer WK, Papadopoulos S, Drumm DA et al. Neurosurgery 2008 Mar;62(3 Suppl 1):201-7; discussion 207-8.
Conclusion : All three registration methods provided comparable distances to the target tissue for surgical procedures.
Anatomic registration as a less costly and more time-efficient registration method for frameless stereotaxy.

## Slide 17: Pediatric patient
Frameless stereotactic procedures in pediatric patients: safety and diagnostic efficacy. Mary G. Parreño & Xiao Bo & Okezie O. Kanu & Shlomi Constantini & Andrew A. Kanner. Childs Nerv Syst (2011) 27:2137–2140
N:21
Age Mean, 8.7 years Range, 0.9– 18.9 years
18 biopsies and 4 catheter placement.
In children below 2 years of age, the head was positioned and draped in a horseshoe holder, and the reference antenna was fixed to the rigid component of the operating table.

## Slide 18
Frameless, pin less stereotactic neurosurgery in children. John F. Reavey Cantwell, Frank J. Bova, David W. Pincus. J Neurosurg (6 Suppl Pediatrics) 104:392–395, 2006.
Six of the nine patients were younger than 2 years of age.
beanbag devices were used

## Slide 19
Pinless frameless electromagnetic image guided neuroendoscopy in children. Jason L. McMillen & Marianne Vonau & Martin J. Wood. Childs Nerv Syst (2010) 26:871–878.
N: 19 ; 14 patients <3 yrs.
StealthStation AxiEM Neuronavigation System (Medtronic, Louisville, CO, USA).
Utilizes a coil to generate an electromagnetic field around the patient’s head.
Patients on jelly head ring or the Mayfield headrest

## Slide 20: Controversy for its use in DBS
Clinical Motor Outcome of Bilateral Sub thalamic Nucleus Deep-Brain Stimulation for Parkinson’s Disease Using Image-Guided Frameless Stereotaxy. Helen Bronte-Stewart, Stephanie Louie, Sara Batya, Jaimie M. Henderson. Neurosurgery 67:1088–1093, 2010.
N: 31
58% improvement from bilateral STN DBS in the UPDRS III.
All motor sub scores improved significantly(P <.01). The mean reduction in medication was 50%.
CONCLUSIONS: Bilateral STN DBS for PD performed by an experienced team using a frameless approach results in outcomes comparable to those reported with the use of the frame-based technique.

## Slide 21
Comparison of Accuracy and Precision between Frame-Based and Frameless Stereotactic Navigation for Deep Brain Stimulation Electrode Implantation. Hjálmar Bjartmarz , Stig Rehncrona. Stereotact Funct Neurosurg 2007; 85:235–242
n : 14.
The vector of deviation was 2.5 +_ 1.4 mm with the frameless technique and 1.2 +_0.6 with the frame-based technique. The differences were statistically significant (p<0.05–0.001).
At clinical follow-ups, tremor reduction was similar irrespective of the implantation technique.

## Slide 22: Electromagnetic
Application of electromagnetic technology to neuronavigation: a revolution in image guided neurosurgery. Caroline HayHurst, PatriCia Byrne, Paul r. Eldridge, Conor l. MalluCCi. J Neurosurg 111:1179–1184, 2009.
N : 150 procedures
The Stealth station Axiem navigation system (Medtronic, Inc.).

## Slide 23
COMPARISON OF MAGNETIC TRACKING AND OPTICAL TRACKING BY SIMULTANEOUS USE OF TWO INDEPENDENT FRAMELESS STEREOTACTIC SYSTEMS. Christopher R.Mascott. Neurosurgery 57[ONS Suppl 3]:ONS-295–ONS-301, 2005.
N : 70 , both system simultaneous used.
Magnetic referencing and tracking was found to be comparable with optical tracking both with regard to calculated and true surgical accuracy.
Interference from metal objects in the magnetic field was seen rarely.

## Slide 24: SonoWand
Neuro-navigation by Intra-operative Three-dimensional Ultrasound: Initial Experience during Brain Tumor Resection. Geirmund Unsgaard, Aage Gronningsaeter, Steinar Ommedal, Torgrim Lie, Thomas Lango. Neurosurgery 50:804–812, 2002

## Slide 25: Intra OP MRI
INDICATIONS OF INTRAOPERATIVE MRI
Supratentorial tumors- low grade glioma
Pituitary tumor.
Epilepsy surgery- side & size localization by functional MRI
Resection control- cytoreductive surgery of large tumor
Intra-ventricular shunt placement
Needle biopsy
Cyst aspiration
Electrode placement – deep brain stimulation
After aneurismal clipping – MR angiography

## Slide 26
Ferromagnetic cardiac implant – heart valve
Pace- maker (non- compatible)
Steel implant fixation for bone fracture
Ferromagnetic spinal implant
Cochlear implant (non- compatible)

## Slide 27: ADVANTAGE OF INTRAOPERATIVE MRI
Improves the accuracy of craniotomy placement
Reduce the size of bone flap
Surgical navigation can be repeatedly upgraded by intraop.
Reduces the chances of neurological deficit in the surgery of eloquent area of brain
Identify any amount of residual tumor at the end of surgery
Identify any residual hematoma in AVM surgery

## Slide 28: Frameless stereotaxy in spine
Preoperative CT based guiding system
Fluoroscopy based guidance system
O arm based system.

## Slide 29
Preoperative CT based guiding system
Advantage | Disadvantage
Easy | Preop images must be compatible with IG
No radiation exposure | Registration process is difficult
 | Not suitable for minimally invasive procedure

## Slide 30
Frameless stereotaxy to facilitate anterolateral thoracolumbar surgery : technique. E. Thomas Chappell, Laura Pare, Matthew O. Dolich, Mohammed Salepour. Neurosurgery 56[ONS Suppl 1]:ONS-110–ONS-116, 2005.

## Slide 31
Fluoroscopy based guidance system
Advantage | Disadvantage
Registration process is obviated | Radiation exposure
 | Depend on quality of images

## Slide 32
O arm based system
Best at present in terms of quality.
Too costly, bulky needs large OT room, radiation risk to staff.

## Slide 33: THANK YOU